## Slide 1: CPAP Machine

## Slide 2: Requirements for CPAP
Oxygen FiO2 21 to 100%
Pressure CPAP 4 to 7 cms
Flow 2 to 7 liter/min
Temperature 37oC
Humidity 100% RH

## Slide 3: Requirements for CPAP
Oxygen Compressed Air
      Compressed Oxygen
      Blender
Pressure Flow of gas and CPAP Generator
Temperature
Humidity
Humidifier

## Slide 5: CPAP Generator
Depth of immersion- 4cm
Patient lungs
Patient interface
Air Oxygen mixture
Circuit
Obstruction
4cm
4cm

## Slide 6: Oxygen
21%
100%
50%
Compressed
Air
Flow meter
Blender
Compressed
Oxygen

## Slide 7: Humidifier
37o c
39o c
37o c
37o c

## Slide 8: Types of CPAP Machines
Bubble CPAP Machine
Flow Driver CPAP Machine
Ventilator CPAP

## Slide 9: Cleaning and Disinfection
Always use disposable circuits
No need to replace circuit routinely
Fill the Humidifier chamber with Distilled water
Use Auto-fill option for filling the chamber
Clean the equipment with a soft cloth
Humidification to avoid water logging

## Slide 10: Troubleshoot
Alarm from the blender
Malfunctioning of Blender
Leaks in the circuit
Condensation in the inspiratory limb
No condensation in the expiratory limb
Inadvertent delivery of high pressures

## Slide 11: Troubleshoot
Blender
Leaks
Safety device

## Slide 12: Condensation and Water

## Slide 13: Summary
Basic Requirements | Pressure Generator Graded Oxygen Warmidification
Pressure | Bubble Generator
Oxygen | Blender Oxygen Compressed Air
Warmidificaiton | Servo Controlled Humidifier