## Slide 1: Oxygen Analyzer

## Slide 2: What is Oxygen analyzer?
An equipment which analyses oxygen concentration being delivered to the neonate
Small cylindrical device consisting of a galvanic cell with sensor
Galvanic sensor has two dissimilar electrodes in aqueous electrolyte KOH

## Slide 3: What is its principle?
Oxygen diffuses from one side of sensor through semi-permeable membrane
Oxidation –reduction reaction generates electric current proportional to oxygen concentration
Current generated is measured and conditioned with external electronics and displayed on digital meter

## Slide 4: How to calibrate?
Switch on the analyzer and expose the sensor to the atmosphere
Displayed value should be 21%
If it does not, control knob is rotated till it reads 21%
The instrument has an accuracy of ± 2%

## Slide 5: How to use?
After calibration, place the sensor inside oxygen hood away from the oxygen inflow end of the tubing
Wait for 5-20 seconds as this is the usual time between placement and detection of the reading (Response time)
Displayed value is the concentration of oxygen

## Slide 6: What precautions should be taken?
Do not pull or bend the sensor
Always calibrate before measurement
Do not leave oxygen monitor stored with power switch on
Avoid high or low temperature to prevent damage to the sensor fluid

## Slide 7: Thank you